## Slide 1: Digitalization of Peristaltic Pump
Presentation on 8th Semester Project ,under the guidance of
Dr. N M Renukappa
Presentation by,
Name | USN
Anirudh G | 4JC12EC018
Darshan A | 4JC12EC037
Mrithyunjaya | 4JC12EC064
M N Sunaada Hebbar | 4JC12EC058

## Slide 2: Advantages of Peristaltic pump
Seal-less design
  Peristaltic pumps have no seals avoiding issues such as leaks of corrosive chemicals and ongoing maintenance
Low maintenance costs
  The only replacement part is the hose or tube a relatively low cost item that can be easily changed in a short time.
Reversible
  Peristaltic pumps are reversible with just change in direction of rotation of motor
Accurate dosing
  The pumps are accurate in dosing; they have a repeatability of ±1%
Dry running and self-priming
  Peristaltic pumps do not require pumped fluid to be continually present pumps can run dry, without costly downtime or repairs.

## Slide 3: Applications
Medicine
  Medical infusion pumps.
  Dialysis machines.
Chemical handling
  Pharmaceutical production.
  Dosing systems for dishwasher and laundry chemicals.
  Chemical treatment in water purification plant.
Testing and research
  Auto Analyzer.
  Analytical chemistry experiments.
BT100L Peristaltic Pump

## Slide 4: Block Diagram of Peristaltic Pump

## Slide 5: Input : 4X4 Numeric Keypad
Numeric keypad has 16 keys arranged in the form of a 4X4 matrix.
The A, B, C, D, * and # keys can be used as navigation keys or special function keys.
It has 8 pins which are connected to Arduino MEGA Digital I/O Pins.

## Slide 6: Output : Graphical LCD Display
128X64 pixel JHD12864E Graphic LCD is used
JHD12864E GLCD
Pin description
The LCD module has 20 pins and it is used in 8 bit parallel mode for interfacing it with Arduino MEGA
LCD features Intensity control of Backlight and contrast control

## Slide 7: Microcontroller Board : Arduino Mega
The Microcontroller controls the rpm of the stepper motor which in turn controls the rate of dispensing the fluid.
The motor driver circuit amplifies the signals from controller to provide the required drive current to the stepper motor.
The Microcontroller generates appropriate signals to control the motor and also provides interface for RS485 communication port which enables connection of other devices to the microcontroller.
Arduino MEGA 2560
4X3 Numeric Keypad and 2 axis Thumb Joystick is interfaced with Arduino MEGA
in order to take input from user.
A 128 X 64 Graphical LCD Display (JHD12864E) is interfaced with Arduino MEGA to provide the required user interface.

## Slide 8: Stepper Motor Driver
BH-MSD-4.5A Micro stepping Driver is used.
Micro stepping driver is particularly suitable
for the applications desired with extremely
low noise and low temperature rise .
Microstepping: 1/2, 1/4, 1/8, 1/16, 1/32,
1/64, 1/128, 1/256
Electrical specification:
Parameter | Minimum | Maximum | Unit
Input voltage(DC) | 20 | 50 | V
Output current | 1.5 | 4.5 | A

## Slide 9: Stepper Motor
BH57SH100 – 4006A manufactured by
Bholanath Engineering Industries is used.
BH57SH100 – 4006A

## Slide 10
Torque Vs RPM Graph of Motor
Wiring Diagram of BH57SH100 – 4006A

## Slide 11: Peristaltic Pump Head
Works based on the principle of alternating compression and relaxation of the tube drawing the fluid contents from one container and dispensing it to the other.
A rotating roller passes along the length of the tube totally compressing it and creating a seal between suction & discharge side of the pump
Support different sized fluid tubes
Many versions of YZ and DG series compatible with BT100L available with
2, 4, 6 and 10 rollers
3 Roller Pump Head
2 Roller Pump Head

## Slide 12: Pump Head and Tube Specification
* WD = Tube Wall Diameter, ID = Tube Inner Diameter

## Slide 13: Features
1) Programmable Configuration
Selection of Pump Head
Selection of Tube Size
2) Bidirectional Operation
Switch to select direction of operation – Forward/Reverse

## Slide 14
3) Calibration
The user can select calibrate option to manually calibrate the flow rate for the selected pump head and tube size
The calibration algorithm is as shown in the Flow chart in the next slide
The user is prompted to calibrate after using the tube many times in order to compensate for the change in flow rate due to expansion of tube.

## Slide 16
4)Volume based mode:
I/p: Volume to be dispensed
O/p: Dispensing the required volume of liquid
Configurable Options:
  Different flow rate dispensing mode – The fluid can be dispensed at Min/Average/Max flow rates (with default average flow rate)
  Start foot switch
5) Time based Mode:
I/p: Volume of Liquid in ml, Time in min.
O/p: Dispensing the required volume of liquid in specified time

## Slide 17
5) Repetitive Dispensing Mode:
I/p: Process Duration in hours , Duration Between Samples in min , Volume of Liquid to be Dispensed for each Sample in ml
O/p: The specified Volume of Liquid is Dispensed between samples at a Default Flow Rate (50% of maximum)
6) Discrete Dispensing Mode:
I/p: Process Duration in hours, Duration Between samples in min, Volume of Liquid to be Dispensed for each sample in ml, Flow rate in ml/min
O/p: The specified Volume of Liquid is Dispensed between samples at the specified Flow Rate

## Slide 18
7) Serial Communication Mode:
I/p: Weight of the Liquid to be Dispensed in grams
A weight balance weighs the liquid dispensed and the reading is sampled by the microcontroller through RS485 Serial Communication Port
O/p: Once the Specified weight of the Liquid is Dispensed the microcontroller stops the liquid flow